# DPP4SiPM command line documentation

The interaction with the DPP4SiPM multichannel analyzer (MCA) is performed through ASCII commands. Either through Ethernet or UART, the same command line applies to setup the device parameters or to retrieve data. The commands requests sent from the host (computer or device that controls the MCA) are always replied by the MCA upon a successful communication establishment.

# 1. ASCII commands format

All the commands share the following structure, as described next.

## Request (from host to MCA)

<Start character> <Command Code> <Parameters and values > <End character>

<Start character>: Always the dollar sign: $ (ASCII 0x24)

<Command code>: Two-letter abbreviation

<Parameters and values>: 32-bit values separated by space character (ASCII 0x20)

<End character>: Carriage return \n (ASCII 0x0D)

Example:

$SP 8 52921 3 \r

- the command $SP means: set parameter

- the 8 means: parameter #8 (see the section MCA commands for details)

- the 52921 means: -0.7699584 in signed Q2.14 fixed-point representation

- the 3 means: binary flag (0b011)

- the \r means: end character, equivalent to ASCII code 0x0D

## Reply (from MCA to host)

The MCA with a similar format, starting with a distinctive reply start character

<Start character> <Received request command> <Values> <End characters>

<Start character>: Always the exclamation mark: ! (ASCII 0x21)

<Received request command>: The same two-character command from the request

<Values>: The payload of the reply. Contents depend on the request command sent.

<End characters>: Two successive characters are returned always: \n\r (ASCII 0x0D 0x0A)

Example:

If the request is $SP 8 52921 3 \r, the reply might be:

$SP\n\r

For the request $RT -1, the reply might be:

!RT 0 0 0 0 10000000 176\n\r

## Error replies

In case the request command is wrong or has not been delivered correctly, an error reply will be sent from the MCA back to the host.

The error reply format is defined as follows:

<Start character> ERROR:<Error code><End characters>

<Start character>: Always the exclamation mark: ! (ASCII 0x21)

ERROR: text followed by the error code

<Error code>: 0 if the request command is not recognized, 1 if the sent parameters are wrong

<End characters>: \n\r (ASCII 0x0D 0x0A)

Example:

If the request is $KR \r, the reply should be:

$ERROR:00\n\r

because KR is an unknown request command

If the request is $SP 8 52921 \r, the reply might be:

$ERROR:01\n\r

because one value is missing to set the parameter #8 (two expected)

# MCA commands

The MCA is configured through request commands. In the same way, data such as the spectrum and timer values can be retrieved using the command line. Such request commands always follow the aforementioned format, and are next.

## Get Version ($GV)

Returns the firmware version of the system.

### Parameters

No parameters are required

### Example

$GV\r

!GV 1.0\n\r

## Get Serial Number ($SN)

Returns the serial number of the MCA.

### Parameters

No parameters are required

### Example

$SN\r

!SN 210328B00000B\n\r

## Get Parameters ($GP)

Returns the values for each of the DPP parameters available in the MCA.

## Read Timers ($RT)

Asdf

## Read Spectrum ($RM)

Asdf

## Load Scope ($LS)

Asdf

## Set Parameters ($SP)

Asdf

## Acquisition Start/Stop ($AQ)

Asdf

## Clear Spectrum ($CS)

Asdf

## Set Serial Number ($SS)

Asdf

## Keep Alive (~~~)

Asdf

# Digital pulse processor

The digital pulse processor (DPP) in the MCA is configurable, aiming at matching the detector characteristics for each use case. The configurable parameters can be retrieved or set by using the request commands $GP and $SP, respectively. These values are grouped according to the internal building blocks of the DPP and should be appended sequentially in unsigned 32-bit format every time the $SP command is sent.

## Pulse shaper – slow

Parameter index: 1

| Setting | Description | Signedness and fixed-point format (Qm.n) | Default value (unsigned 32-bit) | Default value (human readable) |
| --- | --- | --- | --- | --- |
| Tclk | Period of the DPP clock in the FPGA, in nanoseconds (ns). | U – Q32.0 | 20 | 20 |
| Taur | Rising edge constant of the detector pulse shape, in ns. | U – Q32.0 | 220 | 220 |
| Taud | Falling edge (decay time) constant of the detector pulse shape, in ns. | U – Q32.0 | 1145 | 1145 ns |
| Taupk | Peaking time of the trapezoidal shaper, in ns. | U – Q32.0 | 3000 | 3000 ns |
| Taupk_top | Flat-top of the trapezoidal filter, in ns. | U – Q32.0 | 0 | 0 ns |
| Gain | Digital gain at the output of the trapezoidal filter. | U – Q1.25 | 33554432 | 1.0 |
| b10 | exp(-Tclk/Taud) | U – Q0.25 | 32973417 | 0.9826844036579132 |
| na_inv | Tclk/taupk | U – Q0.18 | 1747 | 0.006664276123046875 |
| na | Taupk/Tclk | U – Q32.0 | 147 |  |